## Slide 1: New Frontiers

## Slide 2: Delta-Sigma Beamformer
A delta-sigma A/D compromises output bit density for increased sampling rates.
For a delta-sigma beamformer, delays and sums can be performed on the single bit outputs.
A properly designed delta-sigma beamformer can achieve similar signal-to-noise ratio as a typical multiple bit A/D.

## Slide 3: Delta-Sigma Beamformer
10bitA/D
@40MHz
FIFO
FIFO
interp
interp
10bitA/D
@40MHz
19bits
@40MHz
Conventional Beamformer

## Slide 4: Delta-Sigma Beamformer
A/D
@320MHz
FIFO
FIFO
10bitA/D
@40MHz
21bits
@40MHz
Delta-Sigma Beamformer
Filter/ Decimation
9bits
@320MHz

## Slide 5: Real-Time Volumetric Imaging
Volumetrics Medical Imaging.
Real-time volumetric, 3D imaging.
Cardiology applications.
Simultaneous display of multiple B-mode images and C-scan images.
No surface rendering, volume rendering or shading of data.

## Slide 6: Real-Time Volumetric Imaging
2.5MHz 2D array.
Circularly symmetric 16mm on a side.
Approximately 2000 elements, 512 of which are connected.
256 elements for transmit, 256 elements for receive.

## Slide 7: Real-Time Volumetric Imaging
16 parallel beams on receive.
4096 lines that are spaced approximately 1 degree apart.
Twenty two 64X64 degree pyramidal volumes per second at 15cm.

## Slide 8: Real-Time Volumetric Imaging
Up to 16 views can be displayed simultaneously.
Entire data set is stored. Therefore, various locations and scan planes can be displayed during re-play.
Can be used to increase throughput, and the skill level necessary to produce a high quality scan can be decreased.

## Slide 9: Real-Time Volumetric Imaging
Clinical applications include LV function, RV function, LV volumes, myocardial mass and myocardial perfusion.
No Doppler capabilities.
The future of ultrasonic imaging ?