## Slide 1: Two-Dimensional Transducer Arrays

## Slide 2: Motivations: Elevational Focusing
1D arrays have only fixed elevational focusing capabilities. The elevational focusing quality is determined by a mechanical lens.
2D arrays must be used in order to have electronic, dynamic 3D focusing.

## Slide 3: Motivations: Elevational Focusing
Although not apparent, elevational focusing quality determines the “slice thickness” and is critical in image quality (including contrast resolution, noise characteristics and elevational resolution).
slice thickness

## Slide 4: Motivations: Aberration Correction
Reported correlation lengths of phase and energy level fluctuations from abdominal measurements are around 5mm and 3mm respectively.

## Slide 5: Motivations: Aberration Correction
Severe averaging effects will occur if the elevational size is large.
Measurements have shown that 1D correction alone does not necessarily improve the image quality.
2D arrays are necessary for adequate aperture sampling.

## Slide 6: Motivations: Real-Time 3D Imaging
Current 3D imaging (including flow imaging) is done by reconstruction using a set of 2D images. This is not real-time.
Fully sampled (i.e., spatial Nyquist criterion is satisfied in both directions) arrays must be used to allow full 3D focusing and steering capabilities.

## Slide 7: Complications
For a fully sampled aperture with a typical size, the total channel count may exceed 10,000 - almost impossible for current technologies.
As the element size decreases, electrical impedance significantly increases, thus resulting in poor signal to noise ratio.

## Slide 8: Complications
Inter-connection from each channel to the front-end electronics becomes complicated with 2D arrays.

## Slide 9: Reduction of Channel Count
Sparse array:
  A random selection of elements are removed from the periodic dense array. Only a fraction of the original elements remains.
  Mainlobe is un-affected. Grating lobes are avoided. Sidelobes are higher.
  Electrical impedance is un-changed.
  Difficult to manufacture.

## Slide 10: Reduction of Channel Count
1.5D array:
  Aperture in elevation is under-sampled.
  Electrical impedance is reduced.
  Total channel count is reduced.
  Lack of elevational steering.
  Full 3D focusing quality on the 2D image plane.
  Also known as an-isotropic arrays.

## Slide 11: Sampling of 1.5D Arrays
Uniform sampling.
Fresnel sampling.
Geometric delays are symmetric.

## Slide 12: Sparse Periodic Arrays
Despite of the periodicity, grating lobes are avoided by placing them at different locations.

## Slide 13: Sparse Periodic Arrays
Two-way beam pattern is determined by two-way aperture function, i.e., the convolution of the transmit aperture with the receive aperture.
By carefully choosing the aperture functions, desired response can be synthesized.

## Slide 14: Sparse Periodic Arrays
Only valid for CW and in-focus. 4-to-1 reduction is reasonable.
Signal-to-noise ratio is affected.
In general, this is a synthetic aperture approach.
This approach can be extended to 2D.

## Slide 15: Lowering Impedance
Electrical impedance significantly increases (high resistance and low capacitance) with small transducer area.
For typical system characteristics, large impedance means poor sensitivity.
Impedance can be reduced by many methods, including using multiple layer piezoelectric material.

## Slide 16: Multi-Layer Ceramic
Acoustically in series and electrically in parallel.
single layer
multiple layer

## Slide 17: Multi-Layer Ceramic
On transmit, the acoustic output pressure is increased by a factor of N (number of layers) assuming the same drive voltage.
The output pressure is also increased by matching the impedance.

## Slide 18: Multi-Layer Ceramic
On receive, the received voltage is reduced by a factor of N assuming the same returning pressure.
Capacitance is increased by a factor of N2, due to the parallel connection and smaller distance between plates.
The receive improvement depends on specific situations.

## Slide 19: Multi-Layer Ceramic
Both KLM and finite element models have been used to predict the performance and compared to measurement results.
An alternative method is to have multiple layers on transmit and signal layer on receive.
The coaxial cable capacitance can be reduced by integrating front-end with transducers.

## Slide 20: 2D Aberration Correction
Measurements done by various groups have shown the necessity of 2D arrays for phase aberration correction.

## Slide 21: Complications for Real-Time 3D Imaging
Channel count.
System complexity associated with increased channel count.
The number of beams in a 3D imaging is dramatically increased, thus reducing “frame rate”.
Parallel beam formation is desired.

## Slide 22: Very Large Conformal Arrays
To further enhance the resolution, very large arrays may be used.
Complications with very large arrays:
  Channel count.
  Contact.
  Acoustic access.

## Slide 23: Very Large Conformal Arrays
To maintain adequate contact, the arrays need to be flexible.
The irregular array geometry results in beamforming errors (similar to near field phase aberration).
Distortion due to irregular array geometry can be corrected by correlation based correction methods.

## Slide 24: Very Large Conformal Arrays
With very large apertures, acoustic obstructions due to air pockets and bone structures may not be avoided.
A portion of the array elements may receive no signals or strong reverberations.
Blocked elements also degrade image quality.

## Slide 25: Very Large Conformal Arrays
Degradation from blocked elements can be reduced by using information contained in beams from different directions.
The parallel receive compensation method effectively reduces artifacts regardless of the source.
Very complicated.