## Slide 1: Speckle Correlation Analysis

## Slide 2: Speckle Formation
Speckle results from coherent interference of un-resolvable objects.
sample volume
transducer

## Slide 3: Speckle Formation
In diagnostic ultrasound, the size of tissue micro-structures is often much smaller than a typical wavelength.
Pulse-echo ultrasonic images are formed using the phase information.
Speckle appears as brightness variations and obscure the underlying information.

## Slide 4: Speckle First-Order Statistics
Re
Im

## Slide 5: Speckle First-Order Statistics

## Slide 6: Speckle Second-Order Statistics
The auto-covariance function of the received phase-sensitive signals (i.e., before envelope detection) is simply the convolution of the system’s point spread function if the insonified region is
  macroscopically slow-varying.
  microscopically un-correlated.

## Slide 7: Speckle Second-Order Statistics
The auto-covariance function of post-detection signals also have similar properties.
The shape of a speckle spot (assuming fully developed) is simply determined by the shape of the point spread function.
The higher the spatial resolution, the finer the speckle pattern, and vice versa.

## Slide 8: Speckle Statistics
The above statements do not hold if the object has structures compared to or larger than the ultrasonic wavelength.
Rician distribution is often used for more general scatterer distribution.
Rayleigh distribution is a special case of Rician distribution.

## Slide 9: Speckle and Contrast Resolution
Speckle noise is fundamentally limiting contrast resolution.
Contrast resolution can be represented by the following equation:
I1
I2

## Slide 10: Speckle and Contrast Resolution
Contrast resolution is improved with finer speckle pattern (i.e., higher spatial resolution).
In-coherent processing must be used to reduce speckle variations.
In-coherence can be introduced in many different ways, but all at the price of spatial resolution.

## Slide 11: Speckle Pattern Correlation
Phase of the received signal is determined by both the frequency and the length between the object and the transducer.
To effectively reduce speckle, correlation between two measurements of the same object must be understood.

## Slide 12: Speckle Correlation: Lateral
correlation coefficient
displacement
L/2

## Slide 13: Speckle Correlation: Lateral
Assuming the target is at focus, the correlation roughly decreases linearly as the lateral displacement increases.
The correlation becomes zero when the displacement is about half the aperture size.
There exists some variations between different models.

## Slide 14: Speckle Correlation: Lateral
Correlation may decrease in the presence of non-ideal beam formation.
The choice of lateral displacement determines the correlation between images as well as the number of images to be compounded.
Lateral resolution is degraded.

## Slide 15: Speckle Correlation: Lateral
Assuming there are N un-correlated images, the speckle variations are reduced to 1/N0.5.
“Optimal” scheme can be developed by considering both the degradation in spatial resolution and reduction in speckle variations.
Frame rate may be reduced.

## Slide 16: Speckle Correlation: Elevational
elevation
main image plane

## Slide 17: Speckle Correlation: Elevational
With a two-dimensional array, un-correlated or partially correlated images can be obtained by steering elevationally.
The correlation between images is simply determined by the overlap of the elevational beam pattern.
Only small displacements are required.

## Slide 18: Speckle Correlation: Spectral
f

## Slide 19: Speckle Correlation: Spectral
Similar to lateral compounding, the signal spectrum can be divided into several sub-bands.
The correlation between sub-bands decreases linearly as the overlap decreases, assuming rectangular sub-bands.

## Slide 20: Speckle Correlation: Spectral
Speckle variations are reduced at the price of axial resolution.
“Optimal” scheme can also be developed by considering both spatial and contrast resolution.
Temporal compounding also results in loss in axial resolution.

## Slide 21: van Cittert-Zernike Theorem
A theorem originally developed in statistical optics.
It describes the second-order statistics of the field produced by an in-coherent source.
The insonification of diffuse scatterers is assumed in-coherent.
It is different from the aforementioned lateral displacement.

## Slide 22: van Cittert-Zernike Theorem
The theorem describes the spatial covariance of signals received at two different points in space.
For a point target, the correlation of the two signals should simply be 1.
For speckle, correlation decreases since the received signal changes.

## Slide 23: van Cittert-Zernike Theorem
The theorem assumes that the target is microscopically un-correlated.
The spatial covariance function is the Fourier transform of the radiation pattern at the point of interest.

## Slide 24: van Cittert-Zernike Theorem
radiation pattern
correlation

## Slide 25: van Cittert-Zernike Theorem
The theorem states that the correlation coefficient decreases from 1 to 0 as the distance increases from 0 to full aperture size.
The correlation is independent of the frequency, aperture size, …etc.

## Slide 26: van Cittert-Zernike Theorem
In the presence of tissue inhomogeneities, the covariance function is narrower since the radiation pattern is wider.
The decrease in correlation results in lower accuracy in estimation if signals from different channels are used.

## Slide 27: van Cittert-Zernike Theorem
distance
correlation

## Slide 28: Speckle Tracking
Estimation of displacement is essential in many imaging areas such as Doppler imaging and elasticity imaging.
Speckle targets, which generally are not as ideal as points targets, must be used in many clinical situations.

## Slide 29: Speckle Tracking
From previous analysis on speckle analysis, we found the local speckle patterns simply translate assuming the displacement is small.
Therefore, speckle patterns obtained at two instances are highly correlated and can be used to estimate 2D displacements.

## Slide 30: Speckle Tracking
Displacements can also be found using phase changes (similar to the conventional Doppler technique).
Alternatively, displacements in space can be estimated by using the linear phase shifts in the spatial frequency domain.

## Slide 31: Speckle Tracking
Tracking of the speckle pattern can be used for 2D blood flow imaging. Conventional Doppler imaging can only track axial motion.
Techniques using phase information are still inherently limited by the nature of Doppler shifts.