## Slide 1: Aberration Correction

## Slide 2: Sound Velocity Inhomogeneities
transducer array
v1 v2 v3
point of interest
body wall
viscera

## Slide 3: Sound Velocity Inhomogeneities

## Slide 4: Sound Velocity Inhomogeneities
Sound velocity variations result in arrival time errors.
Most imaging systems assume a constant sound velocity. Therefore, sound velocity variations produce beam formation errors.
The beam formation errors are body type dependent.

## Slide 5: Sound Velocity Inhomogeneities
Due to beam formation errors, mainlobe may be wider and sidelobes may be higher.
Both spatial and contrast resolution are affected.
no errors
with errors

## Slide 6: Near Field Assumption
Assuming the effects of sound velocity inhomogeneities can be modeled as a phase screen at the face of the transducer, beam formation errors can be reduced by correcting the delays between channels.
beam formation
correction
geometric delay
velocity variations
aligned

## Slide 7: Correlation Based Method
Time delay (phase) errors are found by finding the peak of the cross correlation function.
It is applicable to both point and diffuse targets.

## Slide 8: Correlation Based Method
The relative time delays between adjacent channels need to be un-wrapped.
Estimation errors may propagate.

## Slide 9: Correlation Based Method
Two assumptions for diffuse scatterers:
  spatial white noise.
  high correlation (van Cittert-Zernike theorem).
filter
correlator
x

## Slide 10: Correlation Based Method
Correlation using signals from diffuse scatterers under-estimates the phase errors.
The larger the phase errors, the more severe the underestimation.
Iteration is necessary (a stable process).

## Slide 11: Alternative Methods
Correlation based method is equivalent to minimizing the l2 norm. Some alternative methods minimize the l1 norm.
Correlation based method is equivalent to a maximum brightness technique.

## Slide 12: Baseband Method
The formulation is very similar to the correlation technique used in Color Doppler.

## Slide 13: Baseband Method
CORDIC
CORDIC
I
Q
I
Q
acc.
acc.
acc.
Q sign bit
sign control

## Slide 14: One-Dimensional Correction: Problems
Sound velocity inhomogeneities are not restricted to the array direction. Therefore, two-dimensional correction is necessary in most cases.
The near field model may not be correct in some cases.

## Slide 15: Two-Dimensional Correction
Using 1D arrays, time delay errors can only be corrected along the array direction.
The signal received by each channel of a 1D array is an average signal. Hence, estimation accuracy may be reduced if the elevational height is large.
2D correction is necessary.

## Slide 16: Two-Dimensional Correction
Each array element has four adjacent elements.
The correlation path between two array elements can be arbitrary.
The phase error between any two elements should be independent of the correlation path.

## Slide 17: Full 2D Correction
(1,1)
(1,3)
(1,2)
corr
corr
(3,1)
(3,3)
(3,2)
corr
corr
(2,1)
(2,3)
(2,2)
corr
corr
corr
corr
corr
corr
corr
corr

## Slide 18: Row-Sum 2D Correction
(1,1)
(1,3)
(1,2)
corr
corr
(3,1)
(3,3)
(3,2)
corr
corr
(2,1)
(2,3)
(2,2)
corr
corr
corr
corr

## Slide 19: Correlation Based Method: Misc.
Although conceptually simple, correlation based method requires extensive computations.
Limited human studies have shown its efficacy, but the performance is not consistent clinically.
2D arrays are required to improve the 3D resolution.

## Slide 20: Displaced Phase Screen Model
Sound velocity inhomogeneities may be modeled as a phase screen at some distance from the transducer to account for the distributed velocity variations.
The displaced phase screen not only produces time delay errors, it also distorts ultrasonic wavefronts.

## Slide 21: Displaced Phase Screen Model
Received signals need to be “back-propagated” to an “optimal” distance by using the angular spectrum method.
The “optimal” distance is determined by using a similarity factor.
phase screen

## Slide 22: Displaced Phase Screen Model
After the signals are back-propagated, correlation technique is then used to find errors in arrival time.
It is extremely computationally extensive, almost impossible to implement in real-time using current technologies.

## Slide 23: Wavefront Distortion
Measurements on abdominal walls, breasts and chest walls have shown two-dimensional distortion.
The distortion includes time delay errors and amplitude errors (resulting from wavefront distortion).

## Slide 24: Time Reversal Mirror
Due to the reciprocal nature of pulse-echo imaging, perfect focusing can be achieved by transmitting a pulse (p(T-t)) which is the temporal reverse of a previously received signal (p(t)) .
This is essentially a matched filter technique.

## Slide 25: Time Reversal Mirror
A point reflector is required, thus making it non-ideal for diffuse scatterers.
It can be used in lithotripsy for automatic focusing.
time reversed
geometric delay
velocity variations
perfect focusing

## Slide 26: Translating Transmit Aperture
Correlation between signals from a speckle target is determined by the van Cittert Zernike theorem.
In the presence of phase aberrations, correlation may be further reduced.
Alternative transmit-receive aperture arrangement can be used to increase correlation.

## Slide 27: Translating Transmit Aperture
In the absence of phase aberrations, two sets of received signals separated by x is perfectly correlated if the transmit apertures are displaced by - x.
T1/R2
R1/T2

## Slide 28: Phase Conjugation
phase screen at face of transducer
displaced phase screen
f
f
phase
phase

## Slide 29: Phase Conjugation
Simple time delays result in linear phase shift in the frequency domain.
Displaced phase screens result in wavefront distortion, which can be characterized by non-linear phase shift in the frequency domain.

## Slide 30: Phase Conjugation
Non-linear phase shift can be corrected by dividing the spectrum into sub-bands and correct for “time delays” individually.
In the limit when each sub-band is infinitesimally small, it is essentially a phase conjugation technique.

## Slide 31: Compensation Using Parallel Receive Beams
Fixed direction transmit, all direction receive beam formation can be used to estimate the un-wanted sidelobe contributions.

## Slide 32: Compensation Using Parallel Receive Beams
The parallel receive beams are the convolution of transmit insonification with receive beam pattern.

## Slide 33: Compensation Using Parallel Receive Beams
The un-desired sidelobe contributions can be estimated by using de-convolution.
Both the modeling errors and measurement errors need to be included.
The TLS (Total-Least-Squares) model is used.

## Slide 34: Compensation Using Parallel Receive Beams
Sidelobe contribution is reduced regardless of the cause of it. In fact, it was originally proposed for other imaging artifacts.
It can be combined with correlation based method as a two-step process.
It is very computationally demanding.